## Slide 1
II. DESCRIPTION
I. INTRODUCTION
Recording neural signals are important for healthcare and clinical treatment. Owning to low amplitude of important biopotentials such as Electrocardiogram (ECG), a chopper amplifier is employed to reduce the intrinsic noise while amplifying biomedical input signals. However, the chopper amplifier is associated with some disadvantages such as low input impedance, output ripple. Moreover, the electrode offset (EOS) creates DC current at the electrode caused tissue damage due to the finite dc input-impedance (Zin) of the recording front end. Additionally, the amplifier can be saturated by EOS. Thus, removal EOS and boosting input impedance for chopper amplifier are needed.
References
[1] K. A. Ng and Y. P. Xu, “A compact, low input capacitance neural recording amplifier”, IEEE Trans. Biomedical Circuits Syst., vol. 7, no. 5, pp. 610-620, Oct. 2013.
[2] H. Chandrakumar and D. Marković, "An 80-mVpp Linear-Input Range, 1.6-GΩ Input Impedance, Low-Power Chopper Amplifier for Closed-Loop Neural Recording That Is Tolerant to 650-mVpp Common-Mode Interference," in IEEE Journal of Solid-State Circuits, vol. 52, no. 11, pp. 2811-2828, Nov. 2017.
[3] X.T. Pham, N. T. Nguyen , V. T. Nguyen and J-W Lee, “A 0.6 µW Chopper Amplifier Using a Noise-Efficient DC Servo Loop and Squeezed-Inverter Stage for Power-Efficient Biopotential Sensing”, Sensors, vol. 20, no. 7, pp. 2059, April. 2020.
III. CHIP IMPLEMENTATION AND RESULTS
Fig. 1 Schematic of the proposed chopper amplifier
Fig. 1 shows the schematic of the proposed chopper amplifier (CA) included the Squeezed-inverter (SQI) input stage, following is a folded-cascode and common source amplifiers using 0.8 V supply at the middle and output stages, respectively. The closed-loop gain of CA is 40 dB defined by the ratio of the input capacitor of 2 pF and T-network negative feedback capacitor [1] of 20 fF. By process variation, the Gm1 and Gm2 associate the offset voltages VOS1 and VOS2, respectively. The VOS1, which will be created output ripple, is blocked by high pass filtering approach [2]. While the electrode offset (EOS) is chopped to chopping frequency by chopper input then, it is partially mitigated by Cfb1,2 at the virtual ground. This residual offset is amplified and down-chopped by Gm1 and chopper output resulting VEOS,Gm2 is created in front of Gm2. To address VEOS,Gm2 and VOS2, a DC servo loop (DSL) [3] is added. The DSL senses the output offset, then output voltage of DSL (VO,DSL) applies to the body of the transistor input pairs. Thus, current offsets, which is created by VEOS,Gm2 and VOS2, are compensated leading to the frequency response of the proposed amplifier has high pass cut-off frequency at sub-Hz.
The schematic of the SQI using 0.2 V supply is shown in Fig. 2. To operate with low-voltage supply, the SQI is biased by a negative voltage VNEG, which is generated by the negative voltage generator shown in Fig. 3a. The VNEG is applied to MH3,4 through R3,4 to created bias current of 800 nA to reduce noise [3]. The Monte Carlo simulated results of the negative voltage generator is shown in Fig. 3b. Fig. 3c shows the total power consumption of 0.68 µW of the proposed CA. The proposed of CA integrates the auxiliary path (AUX) [2] to boost the input impedance.
The fabricated of the proposed CCIA, which occupies 0.188 mm2, is shown in Fig. 4. In order to measure the performances, the chip is connected to the dynamic analyzer Agilent 35670A, power supplies, an oscilloscope, a signal generator. Fig. 5 shows the measured results of the frequency responses of the amplifier in cases with and without the DSL. Since the DSL uses variable pseudo resistor which is controlled by an external voltage VPR. When VPR is 0.45 V, the high pass corner at low-frequency of the frequency response achieves 0.7 Hz. And in Fig. 5, the measurement results of the power supply rejection ratios (PSRRs) of 80 and 75 dB from 0.2, 0.8 V supply, respectively, as well as the common mode rejection ratio (CMRR) of 105 dB are shown. The measured of the input impedance is shown in Fig. 6. By controlled SW1,2 in the AUX , the changed pre-charged time can be changed 0.5 to 1.5µS, the measured input impedance can be varied from 0.45 MΩ to 1.8 GΩ . This shows that the auxiliary helps well the boosting input impedance for the proposed CCIA. The measured input referred noise (IRN) is shown in Fig. 7. The IRN achieves 1.3 and 1.8 µVrms in cases disabled and enabled both DSL, AUX, respectively. Hence, the proposed amplifier can be achieved the high input impedance during realizing the electrode offset removal and the ripple rejection. The consumption of CCIA is low power of 0.68 µW from 0.2 and 0.8 V supply. With the performances above, the proposed CCIA can compare favorably with the state of the art.
Fig. 4 Fabricated chip
Fig 5. Measured result of the frequency responses
A 680 nW, 1.8 µVrms, 1.8 GΩ Input Impedance Biomedical Amplifier for Neural Acquisition
Xuan Thanh Pham, Chang-Hyung Choi and Jong-Wook Lee
School of Electronics and Information, Kyung Hee University, jwlee@khu.ac.kr
Acknowledgement
The chip fabrication and CAD tools were supported by the IDEC (IC Design Education Center).
Fig. 2 Schematic of Squeezed-inverter input stage
Fig. 3 a. Schematic and b. Monte Carlo simulation of the negative voltage generator; c. Power breakdown of the IA
Fig 6. Measured result of input impedance
Fig 7. Measured input referred noise results
(a) (b) (c)