## Slide 1
A Two-Electrode Bio-Impedance Readout IC with
Complex-Domain Noise-Correlated Baseline Cancellation
Song-I Cheon†, Haidam Choi†, and Minkyu Je
School of Electrical Engineering, Korea Advanced Institute of Science and Technology (KAIST), †Equally Credited Authors
E-mail: 1000thd2@kaist.ac.kr
Introduction
Measurement Results
Proposed Bio-Impedance Readout IC
Conclusion
References
Impedance in
2-Electrode System
BioZ
ZEL
ZEL
Total Impedance
= 2 x ZEL
+BioZ
1. Large Baseline
2. Large Phase-shift
(10 ~ 30 @ 50 kHz)
Measurement Issue
VIN
VIN
t
t
t
t
Noise Delayed
VOUT
VOUT
Baseline level
Vcancel
Square Signal
Pseudo-sine Signal
Resistive BioZ
Complex BioZ
Vn,cancel
Vn,in
Prior Work : IDAC Based Resistive
Noise-Correlated Baselince Cancellation [1]-[3]
Previous BC
Proposed BC
[1] C. S. Park et al., A Baseline-Tracking Single-Channel I/Q Impedance Plethysmogram IC for Neckband-Based Blood Pressure and Cardiovascular Monitoring, ISSCC, Feb. 2022
[2] T. Zhang et al., A 0.6V/0.9V 26.6-to-119.3W ∆Σ-Bbased Bbio-Iimpedance Rreadout IC with 101.9dB SNR and <0.1Hz 1/f corner, ISSCC, Feb. 2021
[3] H. Ha et al., A Bbio-Iimpedance Rreadout IC with Ddigital-Aassisted Bbaseline Ccancellation for Ttwo-Eelectrode Mmeasurement, ISSCC, Feb. 2019
Cannot cannel out BGR & IDAC noise
 DEM Required
Difficult to apply for pseudo-sine signal injection
Cancel out whole TX noise
 DEM is not required
Cannot delay the noise of TX
Pseudo-sine signal injection is available
Same delay for BioZ noise and BC noise
Differential Difference CBIA
w/ Quiet Chopping
Tunable
Reference
Impedance
Human Body
Complex-domain
Noise-correlated BC
Current
Generator
Readout IC
I-CH
Q-CH
Impedance
Tuner
Pseudo-sine/ Square-wave Gen.
Look-up Table
Current Reference
ZBioZ,DC
∆ZBioZ
fIF
2nd-order
NS-SAR
ADC
9b’
9b’
R
C
eI
eQ
ZEL= 18.7 ~ 53.5
@ 50kHz
51kΩ
47nF
50~200Ω
Dout,Q
Dout,I
fCG
VIN,P
VIN,N
VREF,N
VREF,P
TC1
TC2
TI
DD-CBIA
D-TC
Quiet
Chopping
Input Range
Input-dependent Noise
fIF
fIF
Pseudo-sine Noise Cancellation
VBioZ
VTRI
VOUT
VIN-VREF
Differential Difference CBIA
w/ Quiet Chopping
Gm1
Gm2
IDAC
TC
TI
Gm3
Input Range
Pseudo-sine Noise Cancellation
Input-dependent Noise
Input Range
Pseudo-sine
Noise Cancellation
Input-dependent Noise
Digital
JSSC 2022
ISSCC 2022
Prior CBIA
Measurement Setup for Noise Comparison
Dout,I
DCNT,R
DCNT,C
Dout,Q
5b
I-CH
Q-CH
Impedance Tuner
Pseudo-sine / Square-wave Gen.
Proposed BioZ Readout IC
ADC
9b
Ztest
(2) Complex model
(1) Resistive Model
TRI
Noise Density (V/√Hz)
x4.88 Noise Improvement
Ztest= 500
10-6
10-7
10-5
Frequency (Hz)
10-1
100
Noise Density (V/√Hz)
x2.47 Noise Improvement
w/ Real-domain BC (I-CH)
w/ Complex-domain BC (I-CH)
w/ Complex-domain BC (Q-CH)
Ztest= 500
10-6
10-7
10-5
Frequency (Hz)
10-1
100
w/ Real-domain BC (I-CH)
w/ Complex-domain BC (I-CH)
w/ Complex-domain BC (Q-CH)
Enable 2-Electrode configuration w/ pseudo-sine excitation
DD-CBIA w/ quiet chopping enables substraction using pseudo-sine
Tunable Reference Impedance w/ 9b resolution for each magnitude and phase
Improve noise performance by 2.5x and 4.9x for BioZ with phases of 30° and 60°
Parameters | JSSC 22 | ISSCC 22 | ISSCC 21 | JSSC 20 | ISSCC 19 | JSSC 18 | This Work
Technology (μm) | 0.18 | 0.18 | 0.04 | 0.065 | 0.065 | 0.13 | 0.18
Supply (V) | 1.8 | 1.2 | 0.9/0.6 | 0.5 | 1.2 | 1.5/3.5 | 1.2/0.9
Excitation Signal | SQ | SQ | SQ | PS | SQ | PS | PS & SQ
DEM | Required | - | Required | - | Required | - | Not Required
Baseline Cancellation Domain | Complex | No | Complex | No | Real | No | Complex
Noise Correlation Domain | Real | No | Real | No | Real | No | Complex
Input range (kΩ) | 36 | N/A | 20 | 0.126 | 24 | 10 | 37
Current magnitude (μApk) | 5-200 | 5-100 | 5-100 | 4.2 | 5-100 | 5-50 | 3-100
Current frequency (kHz) | 1-215 | 1-1024 | 1-1024 | 20 | 1-1024 | 4-100 | 1-200
Noise density @ 1Hz (mΩ/) | 0.4 (R=200) 4.4 (R=3.6k) | 0.5 (R=200) | 0.52 (R=200) 2.9 (R=2k) | 15.3 (R=100) | 0.4 (R=200) 7.0 (R=2k) | 0.7 (R<50) 9.5 (R=2k) | 1.05 (R=80) 3.89 (R=2.2k)
Max. SNR (w/ 0° load) | 103.5 | 97 | 101.9 | - | 95.7 | 91.4 | 101.5
∆MAX SNR (w/ 30° / 60° load) a | Not reported | Not reported | Not reported | - | Not reported | - | −1.67 / −0.78
Power of CG (μW) | 22.0-404.1 | 14.4-128.4 | 17.8-103.3 | 4.62 | 31.4-154.7 | 56-95 | 25.17-141.21
Power of AFE (μW) | 15.8-24.2 | 15.8 | 8.8-16 | 3.95 | 12.3-28.3 | 89 | 15
Power of ADC (μW) | 44.2 | - |  | - | 6.6 | - | 4.47
FoMSNR b | 150.0 | - | 152.6 | - | 143.5 | 135.8 | 151.0
Area (mm2) | 1.92 | 0.4 | 0.6 | 4.83 | 0.74 | 0.96 | 1.95
Parameters | JSSC 22 | ISSCC 22 | ISSCC 21 | JSSC 20 | ISSCC 19 | JSSC 18 | This Work
Technology (μm) | 0.18 | 0.18 | 0.04 | 0.065 | 0.065 | 0.13 | 0.18
Supply (V) | 1.8 | 1.2 | 0.9/0.6 | 0.5 | 1.2 | 1.5/3.5 | 1.2/0.9
Excitation Signal | SQ | SQ | SQ | PS | SQ | PS | PS & SQ
DEM | Required | - | Required | - | Required | - | Not Required
Baseline Cancellation Domain | Complex | No | Complex | No | Real | No | Complex
Noise Correlation Domain | Real | No | Real | No | Real | No | Complex
Input range (kΩ) | 36 | N/A | 20 | 0.126 | 24 | 10 | 37
Current magnitude (μApk) | 5-200 | 5-100 | 5-100 | 4.2 | 5-100 | 5-50 | 3-100
Current frequency (kHz) | 1-215 | 1-1024 | 1-1024 | 20 | 1-1024 | 4-100 | 1-200
 | 0.4 (R=200) 4.4 (R=3.6k) | 0.5 (R=200) | 0.52 (R=200) 2.9 (R=2k) | 15.3 (R=100) | 0.4 (R=200) 7.0 (R=2k) | 0.7 (R<50) 9.5 (R=2k) | 1.05 (R=80) 3.89 (R=2.2k)
Max. SNR (w/ 0° load) | 103.5 | 97 | 101.9 | - | 95.7 | 91.4 | 101.5
∆MAX SNR (w/ 30° / 60° load) a | Not reported | Not reported | Not reported | - | Not reported | - | −1.67 / −0.78
Power of CG (μW) | 22.0-404.1 | 14.4-128.4 | 17.8-103.3 | 4.62 | 31.4-154.7 | 56-95 | 25.17-141.21
Power of AFE (μW) | 15.8-24.2 | 15.8 | 8.8-16 | 3.95 | 12.3-28.3 | 89 | 15
Power of ADC (μW) | 44.2 | - |  | - | 6.6 | - | 4.47
FoMSNR b | 150.0 | - | 152.6 | - | 143.5 | 135.8 | 151.0
Area (mm2) | 1.92 | 0.4 | 0.6 | 4.83 | 0.74 | 0.96 | 1.95
RDAC
CDAC
R2
R2
(II) TRI Configuration : Complex-Domain
Noise-Correlated BC
(I) TRI Configuration : Real-Domain
Noise-Correlated BC
RDAC
CDAC
R2
R2
b FoMSNR = SNRmax+10log(BW/Power)
where Power is the sum of the max. power of readout, ADC, and the min. power of CG
a SNR = MAX(20log(R/(2RRMS))), RRMS=RN or Rpp/6.6