## Slide 1
A 43.8-to-662.0µW 27.5-to-878.9fps Frame-Rate-Scalable
Duty-Cycled Electrical Impedance Tomography System
Haidam Choi†1, Gichan Yun†1, Sohmyung Ha2 and Minkyu Je1
1KAIST, Daejeon, South Korea, 2New York University Abu Dhabi, Abu Dhabi, UAE, †Equally Credited Authors
Abstract: The proposed duty-cycled EIT system reduces 62.2% and 97.5% of power consumption when operating at 878.9fps and 27.5fps, respectively, from the total average power consumption of 1753.6μW. A MIMO current-balancing IA halves power, area, and input parasitics, while the FIR filter and digital matched filter suppress unwanted harmonics and reduce settling time.
Introduction
Measurement Results – Operation
Fabricated Chip
Overall Frame-rate-scalable Duty-cycled EIT System: The frame rate can be controlled by adjusting the duration of the sleep phase, allowing the frame rate and Pavg to be flexibly adjusted according to the required performance.
Proposed Circuits
Comparison and Performance Summary
Measurement Results – EIT Imaging
 Deep-muscle Information
 Power Consumption
 Measurement Time
[1] H. Wang et al., “Wearable super-resolution muscle–machine interfacing,” Front. Neurosci., vol. 16, Art. no. 1020546, Nov. 2022.
[2] A. Kyu, H. Mao, J. Zhu, M. Goel, and K. Ahuja, “EITPose: Wearable and Practical Electrical Impedance Tomography for Continuous Hand Pose Estimation,” CHI Conference on Human Factors in Computing Systems, pp. 1–10, May 2024.
[3] A. Adler and W. R. B. Lionheart, “Uses and abuses of EIDORS: an extensible softw A. Adler and W. R. B. Lionheart, “Uses and abuses of EIDORS: an extensible software base for EIT,” Physiol. Meas., vol. 27, no. 5, p. S25-S42, Apr. 2006.
References
The proposed EIT system reduce Pavg up to 662.0μW at 878.9fps and 43.8μW at 27.5fps.
Multi-input Multi-output CBIA halve power and area, supporting multi-channel scalability.
Harmonic-canceling filter reduce settling time with on-chip digital I/Q demodulation.
Challenges in portable, real-time interaction
[1]